## Job Cal
On The Job Calibration
UUC Description |  |  |  | ข้อมูลเครื่อง Infusion Pump
Equipment Name |  | Manufacturer |  |  | Model |  | Serial No. |  |  | ID No. | Flow Rate (mL/h) |  | Resolution
Infusion Pump |  | TERUMO |  |  | TE-112 |  | 1304000297 |  |  |  | 5-100 |  | 0
Warm up time : |  | 9 |  | น. | Finish time : |  | 9.3 |  | น.
Reference Standard |  |  |  | ข้อมูลเครื่อง Infusion Device Analyzer
Instrument |  |  |  |  | Manufacturer |  | Model |  |  | Serial No. | ID No. | Cert. No. |  | Due Date
Infusion Device Analyzer |  |  |  |  | FLUKE BIOMEDICAL |  | IDA 4 PLUS |  |  | 13575 | A&T-INFUSION-01 | 16MD576 |  | 2017-05-30
Warm up time : |  | 9 |  | น. | Finish time : |  | 9.3 |  | น.
Received date : |  | 2016-11-10
Calibration Date : |  | 2016-11-10
Cal. Start : |  | 9.3 |  | น. | Cal. Finish : |  | 11 |  | น.
Environment
| Temperature |  |  | Relativ Humidity |  | Air Pressure
Start : | 25.8 |  | oC | 59% |  | - |  | mbar
Finish : | 25.8 |  | oC | 59% |  | - |  | mbar

## PM Sheet
AT PREVENT & CALIBRATION LIMITED PARTNERSHIP
276 Moo 21, Sila, Amphoe Mueang,
Khon Kaen 40000 Tel.0-4334-7362
ใบรายงานผลการบำรุงรักษา
รหัสเครื่องมือ: |  | BP-017-080 |  | โรงพยาบาล: |  |  | รพสต.โพนสว่าง
ชื่อเครื่องมือ: |  | Digital Blood Pressure |  | แผนก: |  |  | -
ยี่ห้อ: |  | microlife |  | PM.วันที่: |  |  | 30 กรกฏาคม 2560
รุ่น: |  | BP3AG1 |  | ครั้งต่อไป: |  |  | -
หมายเลขเครื่อง: |  | 41091778 |  | PM.ครั้งที่: | 1 | จำนวนครั้ง/ปี |  | -
รายละเอียดการตรวจเช็ค |  |  |  | ผลการตรจเช็ค |  |  | หมายเหตุ
|  |  |  | ผ่าน | แก้ไข | ไม่มี/ไม่ทำ
1. โครงสร้าง / ตัวถังของเครื่อง
2. ความมั่นคง / แข็งแรงของที่ติดตั้ง
3. สภาพของผ้าพันแขน(CUFF)
4. อะแดปเตอร์
5. แบตเตอรี่
6.สภาพของลูกยางบีบ
7. หน้าจอแสดงผล (Display Panel)
8. สวิทซ์ เปิด/ปิด (Power ON/OFF)
9. ปุ่มกด/แผงปุ่มกด (Key Switch)
10. ฝาปิด/ฝาครอบ/แผ่นครอบ
11. ป้ายแสดงรายละเอียดต่างๆ
12. ปั๊มลม
13. การตรวจสอบระบบของเครื่องก่อนการใช้งาน
14. การรั่วไหลของลมจากเครื่อง
15. อัตราการเต้นของหัวใจ
16. ทดสอบความแม่นยำของแรงดัน
17. ความแม่นยำของการวัด (60-240 mmHg)
สรุปผลการดูแลรักษา: |  |  | ผ่าน | ต้องแก้ไข |  |  | หยุดใช้งานเครื่อง
สถานะของเครื่อง: |  | พร้อมใช้งาน
บำรุงรักษาโดย: | Sirichock Simsen |  |  | รับทราบโดย: |  | Kiartikhun Intha

## Equipment
Equipment : | Infusion Device Analyzer |  | Cert No.: CAL16-00513
Model : | IDA 4 PLUS
Serial No. : | 13575
ID No. : | A&T-INFUSION-01
Manufacturer : | FLUKE BIOMEDICAL
Temp. : | 25.8 | oC
Humidity : | 59 | %
Calibration date : | 2016-11-11
Result of calibration | Without adjustment
The reported uncertainty of measurement was based on a standard uncertainty multiplied by a coverage
factor k = 2.00, providing a level of confidence of approximately 95 %.

## ข้อมูลดิบ
NIBP | BC
Value | (80/40) |  | (120/80) |  | (190/120)
| Systolic | Diastolic | Systolic | Diastolic | Systolic | Diastolic
1 | 86 | 50 | 119 | 82 | 190 | 124
2 | 80 | 48 | 117 | 81 | 192 | 124
3 | 84 | 52 | 120 | 80 | 190 | 124
UUC Systolic | Average | SD | Uncertainty
Reading |  |  | Type A
(mmHg) | (mmHg) | (mmHg) | (±mmHg)
80 | 83.3333 | 3.0551 | 0.96609
120 | 118.6667 | 1.5275 | 0.48305
190 | 190.6667 | 1.1547 | 0.36515
UUC Diastolic | Average | SD | Uncertainty
Reading |  |  | Type A
(mmHg) | (mmHg) | (mmHg) | (±mmHg)
40 | 50 | 2 | 0.63246
80 | 81 | 1 | 0.31623
120 | 124 | 0 | 0

## Uncer Budget
Uncertainty Budget of Systolic (mmHg) |  |  |  | BC
Calculate at : |  | 80 | mmHg
Symbol | Source of uncertainty |  |  |  | Value | Prob. | Divisor | Ci | ui | vi or | Veff | K95
[6 empty] | Dist. |  |  |  | veff | 1 | 13.97
UC 1 | Repeatability of Measurement |  |  |  | #DIV/0! | N | 2 | 1 | #DIV/0! | 2 | 2 | 4.53
UC 2 | Drift since last calibration |  |  |  | 0 | R | 1.7321 | 1 | 0 | 8 | 3 | 3.31
UC 3 | Standard Error |  |  |  | 0.27 | N | 2 | 1 | 0.135 | 8 | 4 | 2.87
UC 4 | Error of Resolution |  |  |  | 0.5 | T | 1.7321 | 1 | 0.28868 | 8 | 5 | 2.65
| Combined standard uncertainty |  |  |  |  | normal |  |  | #DIV/0! | #DIV/0! | 6 | 2.52
U | Expanded uncertainty |  |  |  |  | k = | 2 |  | #DIV/0! |  | 7 | 2.43
[11 empty] | 8 | 2.37
Calculate at : |  | 120 | mmHg | [7 empty] | 9 | 2.32
Symbol | Source of uncertainty |  |  |  | Value | Prob. | Divisor | Ci | ui | vi or | 10 | 2.28
[6 empty] | Dist. |  |  |  | veff | 11 | 2.25
UC 1 | Repeatability of Measurement |  |  |  | #DIV/0! | N | 2 | 1 | #DIV/0! | 2 | 12 | 2.23
UC 2 | Drift since last calibration |  |  |  | 0 | R | 1.7321 | 1 | 0 | 8 | 13 | 2.21
UC 3 | Standard Error |  |  |  | 0.27 | N | 2 | 1 | 0.135 | 8 | 14 | 2.2
UC 4 | Error of Resolution |  |  |  | 0.5 | T | 1.7321 | 1 | 0.28868 | 8 | 15 | 2.18
| Combined standard uncertainty |  |  |  |  | normal |  |  | #DIV/0! | #DIV/0! | 16 | 2.17
U | Expanded uncertainty |  |  |  |  | k = | 2 |  | #DIV/0! |  | 17 | 2.16
[11 empty] | 18 | 2.15
Calculate at : |  | 190 | mmHg | [7 empty] | 19 | 2.14
Symbol | Source of uncertainty |  |  |  | Value | Prob. | Divisor | Ci | ui | vi or | 20 | 2.13
[6 empty] | Dist. |  |  |  | veff | 25 | 2.11
UC 1 | Repeatability of Measurement |  |  |  | #DIV/0! | N | 2 | 1 | #DIV/0! | 2 | 30 | 2.09
UC 2 | Drift since last calibration |  |  |  | 0 | R | 1.7321 | 1 | 0 | 8 | 35 | 2.07
UC 3 | Standard Error |  |  |  | 0.27 | N | 2 | 1 | 0.135 | 8 | 40 | 2.06
UC 4 | Error of Resolution |  |  |  | 0.5 | T | 1.7321 | 1 | 0.28868 | 8 | 45 | 2.06
| Combined standard uncertainty |  |  |  |  | normal |  |  | #DIV/0! | #DIV/0! | 50 | 2.05
U | Expanded uncertainty |  |  |  |  | k = | 2 |  | #DIV/0! |  | 100 | 2.025
[11 empty] | >100 | 2
Uncertainty Budget of Diastolic (mmHg)
Calculate at : |  | 40 | mmHg
Symbol | Source of uncertainty |  |  |  | Value | Prob. | Divisor | Ci | ui | vi or
[6 empty] | Dist. |  |  |  | veff
UC 1 | Repeatability of Measurement |  |  |  | #DIV/0! | N | 2 | 1 | #DIV/0! | 2
UC 2 | Drift since last calibration |  |  |  | 0 | R | 1.7321 | 1 | 0 | 8
UC 3 | Standard Error |  |  |  | 0.27 | N | 2 | 1 | 0.135 | 8
UC 4 | Error of Resolution |  |  |  | 0.5 | T | 1.7321 | 1 | 0.28868 | 8
| Combined standard uncertainty |  |  |  |  | normal |  |  | #DIV/0! | #DIV/0!
U | Expanded uncertainty |  |  |  |  | k = | 2 |  | #DIV/0!
Calculate at : |  | 80 | mmHg
Symbol | Source of uncertainty |  |  |  | Value | Prob. | Divisor | Ci | ui | vi or
[6 empty] | Dist. |  |  |  | veff
UC 1 | Repeatability of Measurement |  |  |  | #DIV/0! | N | 2 | 1 | #DIV/0! | 2
UC 2 | Drift since last calibration |  |  |  | 0 | R | 1.7321 | 1 | 0 | 8
UC 3 | Standard Error |  |  |  | 0.27 | N | 2 | 1 | 0.135 | 8
UC 4 | Error of Resolution |  |  |  | 0.5 | T | 1.7321 | 1 | 0.28868 | 8
| Combined standard uncertainty |  |  |  |  | normal |  |  | #DIV/0! | #DIV/0!
U | Expanded uncertainty |  |  |  |  | k = | 2 |  | #DIV/0!
Calculate at : |  | 120 | mmHg
Symbol | Source of uncertainty |  |  |  | Value | Prob. | Divisor | Ci | ui | vi or
[6 empty] | Dist. |  |  |  | veff
UC 1 | Repeatability of Measurement |  |  |  | #DIV/0! | N | 2 | 1 | #DIV/0! | 2
UC 2 | Drift since last calibration |  |  |  | 0 | R | 1.7321 | 1 | 0 | 8
UC 3 | Standard Error |  |  |  | 0.27 | N | 2 | 1 | 0.135 | 8
UC 4 | Error of Resolution |  |  |  | 0.5 | T | 1.7321 | 1 | 0.28868 | 8
| Combined standard uncertainty |  |  |  |  | normal |  |  | #DIV/0! | #DIV/0!
U | Expanded uncertainty |  |  |  |  | k = | 2 |  | #DIV/0!
Calculate at : |  | 300 | mmHg
Symbol | Source of uncertainty |  |  |  | Value | Prob. | Divisor | Ci | ui | vi or
[6 empty] | Dist. |  |  |  | veff
UC 1 | Repeatability of Measurement |  |  |  | 0.0074764 | N | 2 | 1 | 0.0037382 | 35
UC 2 | Drift since last calibration |  |  |  | 0 | R | 1.7321 | 1 | 0 | 8
UC 3 | Standard Error |  |  |  | 0.27 | N | 2 | 1 | 0.135 | 8
UC 4 | Error of Resolution |  |  |  | 0.1 | T | 1.7321 | 1 | 0.057735 | 8
| Combined standard uncertainty |  |  |  |  | normal |  |  | 0.14688 | 83406558.4572
U | Expanded uncertainty |  |  |  |  | k = | 2 |  | 0.29375
Calculate at : |  | 350 | mmHg
Symbol | Source of uncertainty |  |  |  | Value | Prob. | Divisor | Ci | ui | vi or
[6 empty] | Dist. |  |  |  | veff
UC 1 | Repeatability of Measurement |  |  |  | 0.0074764 | N | 2 | 1 | 0.0037382 | 35
UC 2 | Drift since last calibration |  |  |  | 0 | R | 1.7321 | 1 | 0 | 8
UC 3 | Standard Error |  |  |  | 0.27 | N | 2 | 1 | 0.135 | 8
UC 4 | Error of Resolution |  |  |  | 0.1 | T | 1.7321 | 1 | 0.057735 | 8
| Combined standard uncertainty |  |  |  |  | normal |  |  | 0.14688 | 83406558.4572
U | Expanded uncertainty |  |  |  |  | k = | 2 |  | 0.29375
Calculate at : |  | 400 | mmHg
Symbol | Source of uncertainty |  |  |  | Value | Prob. | Divisor | Ci | ui | vi or
[6 empty] | Dist. |  |  |  | veff
UC 1 | Repeatability of Measurement |  |  |  | 0.0074764 | N | 2 | 1 | 0.0037382 | 35
UC 2 | Drift since last calibration |  |  |  | 0 | R | 1.7321 | 1 | 0 | 8
UC 3 | Standard Error |  |  |  | 0.27 | N | 2 | 1 | 0.135 | 8
UC 4 | Error of Resolution |  |  |  | 0.1 | T | 1.7321 | 1 | 0.057735 | 8
| Combined standard uncertainty |  |  |  |  | normal |  |  | 0.14688 | 83406558.4572
U | Expanded uncertainty |  |  |  |  | k = | 2 |  | 0.29375
Calculate at : |  | 450 | mmHg
Symbol | Source of uncertainty |  |  |  | Value | Prob. | Divisor | Ci | ui | vi or
[6 empty] | Dist. |  |  |  | veff
UC 1 | Repeatability of Measurement |  |  |  | 0.0074764 | N | 2 | 1 | 0.0037382 | 35
UC 2 | Drift since last calibration |  |  |  | 0 | R | 1.7321 | 1 | 0 | 8
UC 3 | Standard Error |  |  |  | 0.27 | N | 2 | 1 | 0.135 | 8
UC 4 | Error of Resolution |  |  |  | 0.1 | T | 1.7321 | 1 | 0.057735 | 8
| Combined standard uncertainty |  |  |  |  | normal |  |  | 0.14688 | 83406558.4572
U | Expanded uncertainty |  |  |  |  | k = | 2 |  | 0.29375
Calculate at : |  | 497 | mmHg
Symbol | Source of uncertainty |  |  |  | Value | Prob. | Divisor | Ci | ui | vi or
[6 empty] | Dist. |  |  |  | veff
UC 1 | Repeatability of Measurement |  |  |  | 0.0074764 | N | 2 | 1 | 0.0037382 | 35
UC 2 | Drift since last calibration |  |  |  | 0 | R | 1.7321 | 1 | 0 | 8
UC 3 | Standard Error |  |  |  | 0.27 | N | 2 | 1 | 0.135 | 8
UC 4 | Error of Resolution |  |  |  | 0.1 | T | 1.7321 | 1 | 0.057735 | 8
| Combined standard uncertainty |  |  |  |  | normal |  |  | 0.14688 | 83406558.4572
U | Expanded uncertainty |  |  |  |  | k = | 2 |  | 0.29375
| 3 min | 5 min | 7 min
1 | 12910 | 12911 | 12892
2 | 12911 | 12920 | 12896
3 | 12910 | 12919 | 12897
4 | 12910 | 12918 | 12898
5 | 12910 | 12917 | 12899
6 | 12908 | 12916 | 12898
7 | 12909 | 12919 | 12897
8 | 12909 | 12920 | 12896
9 | 12910 | 12920 | 12895
10 | 12911 | 12919 | 12895
11 | 12911 | 12918 | 12896
12 | 12911 | 12915 | 12897
13 | 12912 | 12913 | 12899
14 | 12912 | 12910 | 12900
15 | 12911 | 12907 | 12899
16 | 12912 | 12904 | 12899
17 | 12913 | 12903 | 12899
18 | 12913 | 12902 | 12898
19 | 12913 | 12901 | 12896
20 | 12915 | 12900 | 12896
21 | 12913 | 12899 | 12896
22 | 12912 | 12901 | 12896
23 | 12912 | 12902 | 12897
24 | 12941 | 12903 | 12896
25 | 12932 | 12903 | 12896
26 | 12926 | 12903 | 12896
27 | 12922 | 12903 | 12896
28 | 12921 | 12903 | 12896
29 | 12920 | 12903 | 12896
30 | 12919 | 12903 | 12898
31 | 12919 | 12902 | 12897
32 | 12920 | 12903 | 12898
33 | 12919 | 12902 | 12897
34 | 12920 | 12901 | 12898
35 | 12919 | 12902 | 12898
36 | 12919 | 12902 | 12897
| 12915.4167 | 12907.9722 | 12911.6944
| 6.9831 | 7.5156 | 1.5296
| 1.1639 | 1.2526 | 0.25493

## ข้อมูลดิบหลัก
Data of Calibration
Department : | โรงพยาบาลเปือยน้อย | Humidity : | 65 | Cert. No. : | 10195
Addresse : | ขอนแก่น | Temparature : | 25.4 | ID. No : | BPD-TTM-001
Equipment : | Digital Blood Pressure | Warm up time: |  | Serial No : | -
Manufacture : | Union Technology | Section : | แพทย์แผนไทย | Test Date : | 10 July 2023
Model : | U82 RH | Finish: |  | Resolution:
|  |  |  | Approve Date : |  |  | 1
NIBP | BC
Value
| Systolic | Diastolic | Systolic | Diastolic | Systolic | Diastolic
1
2
3
UUC Systolic | Average | SD | Uncertainty
Reading |  |  | Type A
(mmHg) | (mmHg) | (mmHg) | (±mmHg)
#VALUE! | #DIV/0! | #DIV/0! | #DIV/0!
#VALUE! | #DIV/0! | #DIV/0! | #DIV/0!
#VALUE! | #DIV/0! | #DIV/0! | #DIV/0!
UUC Diastolic | Average | SD | Uncertainty
Reading |  |  | Type A
(mmHg) | (mmHg) | (mmHg) | (±mmHg)
#VALUE! | #DIV/0! | #DIV/0! | #DIV/0!
#VALUE! | #DIV/0! | #DIV/0! | #DIV/0!
#VALUE! | #DIV/0! | #DIV/0! | #DIV/0!

## Cer Rusult
AT PREVENT & CALIBRATION LIMITED PARTNERSHIP
276 Moo 21, Sila, Amphoe Mueang,
Khon Kaen 40000 Tel.0-4334-7362
Report of Measurement Result
Cert. No. : 23ATP10195
Page 1 of 2
Equipment : Digital Blood Pressure |  |  |  | Department :โรงพยาบาลเปือยน้อย
Manufacture : Union Technology |  |  |  | Addresse :ขอนแก่น
Model :U82 RH |  |  |  | Section : แพทย์แผนไทย
Serial No : - |  |  |  | Temparature : 25.4 °C
ID. No : BPD-TTM-001 |  |  |  | Humidity : 65%
Test Date : 10 July 2023 |  |  |  | Approve Date :
Measurement Result
NIBP
Nominal Value | STD Reading | UUT Reading | Error | Uncertainty | Result
Systolic | Systolic | Systolic | Systolic
(mmHg) | (mmHg) | (mmHg) | (mmHg) | (±mmHg)
80 | 80 | #DIV/0! | #DIV/0! | #DIV/0! | PASS
120 | 120 | #DIV/0! | #DIV/0! | #DIV/0! | PASS
190 | 190 | #DIV/0! | #DIV/0! | #DIV/0! | PASS
Maximum Permissible Error (MPE) : ±8 (mmHg)
Nominal Value | STD Reading | UUT Reading | Error | Uncertainty | Result
Diastolic | Diastolic | Diastolic | Diastolic
(mmHg) | (mmHg) | (mmHg) | (mmHg) | (±mmHg)
40 | 40 | #DIV/0! | #DIV/0! | #DIV/0! | PASS
80 | 80 | #DIV/0! | #DIV/0! | #DIV/0! | PASS
120 | 120 | #DIV/0! | #DIV/0! | #DIV/0! | PASS
Maximum Permissible Error (MPE) : ±8 (mmHg)
Procedure of Test : Test were conducted using in-house test procedure AT-03.
Uncertainty of Measurement
The reported uncertainty is base on a standard uncertainty mutiplied by a converage factor k=2, providing a level of
confidence of approximately 95%
Cert. No. : 23ATP10195
Page 2 of 2
Test Standard Used
Manufacture | Model | S/N |  | Cal. Date | Cert .No.
BC Biomedical | NIBP-1030 | 735AE1524Q |  | 17 February 2023 | MP23-1204
1. This result of test was found accurate as shown on date and place of test only.
2. This result of test was made on requested at the point specified by customer.
3. This test is traceabla to the International System of Units, through :-
- National Institute of Metrology (Thailand).
- National Institute of Metrology (Thailand), through Technology Promotion Association (Thailand-Japan).
4. This certification were carried out using equipment whose measured values are traceable to National Standards, where
these exist.
Test by : ( √ ) Ms.Siriporn Leethongdee |  |  |  | Approved by : _____________________
( ) Mr.Wittawat Khaeonnoi |  |  |  | Mr.Thanapon Chalao